## Slide 1: Introduction to Optically Pumped Magnetometers (OPMs)
Peter D. D. Schwindt

## Slide 2: Properties of a Neutral Atoms
Quantized energy levels
p+
n = 1, l = 0
n = 2, l = 0,1
n = 3, l = 0,1,2
e-
Photon

## Slide 3: Properties of a Neutral Atoms
Spin and magnetic moment
  Electron spin
  Nuclear spin
p+
e-, ½
n = 1, l = 0
n = 2, l = 0,1
n = 3, l = 0,1,2
s = ½
s = -½
n = 1,
l = 0
m = -1
m = 1
F = 1
F = 0
m = 0
Zeeman effect
B
, ½

## Slide 4: Atoms Used in OPMs

## Slide 5: The Rubidium-87 Ground State
Clock—measure hyperfine frequency
Magnetometer—measure Zeeman splitting
Gyromagnetic Ratio = 2p 7 Hz / nT
Atom Energy
mF
-2 -1 0 +1 +2
Hyperfine
splitting
Zeeman splitting
ћwHF
ћwL = 1/4mBB
F = 2
F = 1
F = I + S
I = 3/2, S = 1/2

## Slide 6: Bloch Equations for an OPM
  Atoms diffuse through inert buffer gas (limits atom-wall depolarization)
  Atoms precess under influence of magnetic field at rate given by the gyro magnetic ratio
  Electrons are oriented along circularly polarized pump beam (photon spin s) via optical pumping at rate
  Optical pumping causes depolarization of the collective magnetic moment
  Collisions cause atomic depolarization () and decoherence ()
  Probing the atoms causes depolarization of the collective magnetic moment at a rate
By Lucas Vieira - Own work, Public Domain, https://commons.wikimedia.org/w/index.php?curid=1528090

## Slide 7: Zero-Field/SERF Magnetometer (Vector)
Alkali Vapor Cell
Randomly oriented atomic spins
Optical pumping
Spins align with the pump beam
Pump beam
Circular
polarization
Apply Small Magnetic Field
Spins precess due to magnetic field
B
Out of
plane
Detect with probe beam
Oriented spins rotate the polarization (Faraday rotation)
B
Input polarization
Output polarization
Probe beam

## Slide 8: Signals to Detect
Detect the Pump
Spin Polarization Bloch Equation
Bx = Bz = 0 Steady State Solution
B
Probe beam
Transverse Pump
B0+B1sin(wt)
Atomic Polarization, Pz
Pump Transmission
Atomic Polarization, Px
or Angle of Light Polarization
x
y
z

## Slide 9: Zero-Field SERF Magnetometer

## Slide 10: Two-color pump/probe scheme
Optically pump on D1
Circularly polarized D1 pump beam
D1
795 nm
52S1/2
52P1/2
52P3/2
D2
780 nm
87Rb Fine Structure
Detect on D2
Linearly polarized 780 nm probe beam
Oriented spins rotate the polarization
Input polarization
Output polarization
Two optical resonances in Rubidium (fine structure)
  Use D1 for optical pumping and D2 for probing
  Based on: V. Shah and M. V. Romalis, PRA 80, 013416 (2009)
Dichroic optical components
To analyzer
D1 λ/4
D2 λ/2
Multiorder
Waveplate
780 nm
bandpass filter
Field modulation/phase sensitive detection
required to extract dispersive lineshape
To analyzer
D1 λ/4
D2 λ/2
780 nm
bandpass filter
B0+B1sin(wt)

## Slide 11: Magnetometer Signals
Blue trace: Magnetometer (lock-in) signal

## Slide 12: Linearity of OPMs
Bx
By
From Svenja Knappe

## Slide 13: What is SERF? Spin Exchange Relaxation Free
T2 often limited by spin exchange collisions.
Near zero field and at high density, spin exchange rate >> spin precession frequency
  More F = 2 states, so on average atoms precess in F = 2 direction
  Precession rate is slowed down by nuclear spin and incomplete optical pumping
Record sensitivity: 160 aT / Hz1/2
  Atom shot noise: ~0.1 fT / Hz1/2
SERF in a ~6 mm3 scale cell: 5 fT / Hz1/2
sSE = 2 x 10-14 cm2
sSD = 9 x 10-18 cm2
Rubidium relaxation cross sections
H. D. Dang, A. C. Maloof, M. V. Romalis, Applied Physics Letters 97, 151110 (2010)
W. C. Griffith, S. Knappe, and J. Kitching, Optics Express, 18, 27167-27172 (2010).
F = 2
F = 1
mF
-2 -1 0 +1 +2
87Rb Ground States
Twarge, CC BY-SA 4.0, https://creativecommons.org/licenses/by-sa/4.0, via Wikimedia Commons
g = gyromagnetic ratio
= transverse coherence time
N = number of atoms
t = measurement time

## Slide 14: OPMs for MEG: Early work was all SERF
Princeton/Twinleaf OPM MEG System
2006
NIST Miniature Magnetometer 2012
-now FieldLine Inc.
Sandia Atomic
Magnetometer
2010
QuSpin Inc. Magnetometer
2013

## Slide 15: Emerging Companies Using Zero-Field/SERF OPM
*
*Assuming this is a zero-field/SERF OPM

## Slide 16: Other types of OPM
Unshielded MEG using scalar OPM
Helium-4 OPM
Sensitivity: 50 fT/rt-Hz
Closed-loop, triaxial operation
No cell heating: 10 mW RF discharge
Gradiometric Sensitivity: 15 fT/cm/rt-Hz
Measure field component parallel to ambient field
No crosstalk / magnetically silent
William Fourcault, et al., Opt. Express 29, 14467-14475 (2021)
Limes, M.E., et al., Phys. Rev. Appl. 14, 011002 (2020)

## Slide 17: Helium-4 OPM - principle
photo-detector
pump/probe
laser beam
Parametric resonances
along each direction
with 2 rf fields
and 1 laser beam
Signal (Arb. Units)
Magnetic field amplitude (nT)
0
50
100
-100
-50
-1
-2
-3
demodulation
F. Beato et al. Phys. Rev. A 98, 053431 (2018)
Linear polarization
Magnetic ALIGNED
state
6 cm
2x2 cm

## Slide 18: Pulsed Pump Magnetometer
Operate at nonzero background field
Extraordinary sensitivity without magnetic shielding
Free-induction decay frequency
PUMP PULSE
PUMP PULSE
High performance Optical Magnetic Gradiometers
Zero field SERF OPM
By
σ+ Light
87Rb Spin
θ∝B
Measure
spin angle
Measure
spin frequency
87Rb Spin
Probe
By
87Rb Spin
σ+ Pump
Pump
<1 µs
Probe
~1 ms
f = ɣB

## Slide 19: SQUID vs OPM
SQUID
SERF OPM
Bx
By
Integrates over volume of laser beam
Sensitive axis determined by field modulation (Some SERF designs use laser beam geometry to define sensitivity axis.)
Integrates over the area of pickup loop
Sensitive axis determined by loop geometry
http://hyperphysics.phy-astr.gsu.edu/hbase/Solids/Squid.html

## Slide 20: SQUID vs OPM
 | Sensitivity | Bandwidth | Sensitive axis | Dynamic range (Max field)
Zero field, SERF OPM | 5-20 fT/rt-Hz | 80-150 Hz (open loop) 350 Hz (closed loop)** | 1, 2, or 3 axis (3rd axis from 2nd OPM) | ±2 nT (80 nT) open loop* ±150 nT (150 nT) closed loop**
He OPM | 50 fT/rt-Hz | 2 kHz (open loop) 1.35 kHz (closed loop) | 3 axis (3rd axis 6x worse sensitivity) | ±300 nT (unknown)
Scalar OPM gradio-meter—Unshielded | 15 fT/cm/ rt-Hz | 300 Samples/s (adjustable) | Parallel to am-bient bias field | Unknown (Earth’s field)
SQUID | 2-3 fT/rt-Hz | 1–5 kS/s (adjustable) | Normal to pickup loop | ±20 nT (Unshielded operation possible)
*Quspin Neuro-1: Neuro-1 flyer
**Fieldline: Alem O, et al.,Front. Neurosci. 17, 1190310 (2023)
Unknown = Unknown to this author

## Slide 21: SERF OPMs: Formulating Gain/Sense-Angle Error
Cross Axis Projection Error (CAPE) Effects
Rotation of the sensitive axis
Reduction of gain
Phase errors
  Amir Borna, Joonas Iivanainen, Tony R. Carter, Jim McKay, Samu Taulu, Julia Stephen, Peter D.D. Schwindt, “Cross-Axis projection error in optically pumped magnetometers and its implication for magnetoencephalography systems,” NeuroImage, 247, 118818 (2022).

## Slide 22: Sensitive Axis Rotation: Simulation vs. Measurement
CAPE rotates the sensitive axis by 2.86 °/nT (measured) and 3.33 °/nT (simulated).
The slight difference between the two is attributed to different relaxation rates.

## Slide 23: Mitigation of CAPE
Zero the fields (easier)
  Zero the field at each sensor (< nT)
  Zero the field in the shield
Closed-loop, 3-axis OPM (harder)
  Upside: It’s a 3-axis measurement!
    Can we maintain good sensitivity?
  Upside: Fixed gain. Sensors are always calibrated.
  Downside: cross-talk between sensors.
QuSpin Gen 3 QZFM
Triax Variant: <23 fT/√Hz in 3-100 Hz band (typical 15 fT/√Hz all axis simultaneous)
Uses 2 perpendicular beams
Sandia Person-Sized Shield Coils
Nottingham Biplanar Coils
NeuroImage, 181, 760-774 (2018)
NeuroImage, 236, 118025 (2021)
NeuroImage, 252, 119027 (2022)

## Slide 24: Interference and crosstalk
Interference
  Leakage of the modulation field from one sensor to another
  Can easily be ameliorated through sensor calibration
Crosstalk in closed loop system
  Cancelling field at one sensor is sensed by another sensor
  Mitigation: Design feedback coils that minimize cross talk.
  Mitigation: A cross-talk matrix may need to be measured for each sensor configuration.

## Slide 25: Reporting results
Sensitivity results should be normalized by the bandwidth.

## Slide 26: Reporting results
Table Vibration
Gradiometrically inferred sensitivity = (Ch1 – Ch2)/
Gradiometer sensitivity = (Ch1 – Ch2)/1.8 cm
1.8 cm = gradiometer baseline

## Slide 27: Conclusions
OPMs are fundamentally different than SQUID sensors
  OPMs integrate field over a volume, not an area
  Bandwidth reduced compared to SQUID, at least for SERF OPMs
  Sensitivity of today’s OPM are reduced compared to SQUIDs
    Mitigated by being closer to the head
    Prospects for better OPM sensitivity: Best OPM 0.16 fT/rt-Hz
  OPM respond to all components of the field
    “Vectorization” methods needed to read out specific field components
    Three-axis detection possible
    Systematic errors arise from offset field
SERF and He OPM need a low-field environment
Scalar gradiometer operates in Earth’s field
T ~ 180 C
~ 10 mm
[1]
[1]-modified
~ 15 mm

## Slide 28: Back up slides

## Slide 29: Optical Pumping
2P1/2
Lifetime ~ 27 ns
2S1/2
mS = -½
mS = ½
s+
2P1/2
2S1/2
mS = -½
mS = ½
s+

## Slide 30: Hyperfine Ground State
Ground State Hamiltonian:
  Spin angular momentum
  Nuclear angular momentum
  Total angular momentum
  , (87Rb), (133Cs)
  87Rb : GHz,
  133Cs: GHz,
At low field the nuclear and electron spins are combined,
  ,
  ,
  2.8 MHz/G or 28 Hz/nT (gyromagnetic ratio)
  87Rb: 700 kHz/G or 7 Hz/nT
  133Cs: 350 kHz/G or 3.5 Hz/nT

## Slide 31: Sensitivity Limits
Atom shot noise limit
Decoherence limits noise to 1/N1/2 scaling
Sensitivity improved by increasing T2 or N
Most AMs do not operate at this limit.
Photon shot noise limit
Most AMs operate at or near the photon shot noise limit
Probe intensity contributes to T2 so Nph cannot be made arbitrarily high.
Magnetic
Sensitivity
g = gyromagnetic ratio
T2 = transverse coherence time
N = number of atoms
t = measurement time

## Slide 32: Simulation vs. Measurement, CAPE
OPM’s cross-axis projection error for a 25 Hz, sinusoidal signal with an amplitude of 0.33 nT.

## Slide 33: Impact of CAPE on MEG Signals: M20 Median Nerve Response
Wide bandwidth of SEF response at 20 ms (M20) reveals possible effects of CAPE.

## Slide 34: Impact of CAPE on OPM-MEG Systems’ Localization Capability
The mean and standard deviation (bars) for source localization error; the x-axis is the standard deviation of the remnant static magnetic field on the laser propagation axis selected from a normal gaussian distribution.
Poster Number: IM-29
Amir Borna

## Slide 35: On-Sensor coils for the Gen 2 Sensor
Designed by Joonas Iivanainen with bfieldtools
Coils:
  X: 2 coils
  Y: 2 coils
  Z: 4 coils
Channel
Locations